BROCK UNIVERSITY

RADIATION SAFETY POLICY AND PROCEDURES

Radiation Safety Committee

Brock University

Revised

May 2012

CONTENTS	                                                               Page

A.     RADIATION SAFETY - ADMINISTRATION	3

1.     Radiation Safety Committee (RSC)	3

2.     Radiation Safety Officer (RSO)	4

3.     Internal Permit Holders and Responsibilities	5

4.     Radiation User and Responsibilities	7

B.     RADIATION - WORK PROCEDURES	8

1.	Definition of a Nuclear Energy Worker	8

2.	Medical Examinations	8

3.	Exemption Quantity of a Radioisotope	8

4.	Permissible Exposures	8

5.	Personal Monitoring	9

6.     Laboratory Monitoring	9

7.	Decontamination Procedures	11

8.	Authorization for Use of Radioisotopes	11

9.	Procurement of Radioisotopes	12

10. 	Use and Storage of Radioisotopes	12

11.	Termination of Radioisotope Operation: Decommissioning Rooms	12

C.     RADIOISOTOPE – Safe Handling, Waste Storage and Disposal	13

1.    Responsibility of Permit Holder in Handling and Disposal	13

2.    Responsibility of Radioisotope Users in Handling and Disposal	13

3.	Containment	14

4.	Disposal of Level One Radioactive Wastes	14

5.	Disposal of Level Two Radioactive Wastes	15

D.     Radioisotopes- General Requirements and Precautions	15

E.     Radiation Signs and Labels	15

1.	 Radioisotope Laboratories	15

2.	Other Radiation Areas	15

3.     Unauthorized Use of Radiation-Warning Signs	16

F.	 X-ray And other Radiation Producing Devices - Work Procedures	16

G.     Radiation Incidents - Emergency Procedures	16

APPENDIX 1: Emergency Procedures Flowchart	17

APPENDIX 2: Inspection Report Of Radioisotope Facilities	19

APPENDIX 3: Radioisotope Use And Disposal	21

APPENDIX 4: Lab Specific Training In Radioisotope Use	23

APPENDIX 5: Maximum Permissible Levels	24

APPENDIX 6: Permissible Activity In Radioisotope Laboratory	26

APPENDIX 7: Classification and Exemption quantity of Radioisotopes	27

APPENDIX 8: Operation and Termination Permits	37

APPENDIX 9: References	44

# A.     RADIATION SAFETY - ADMINISTRATION

1.     Radiation Safety Committee (RSC)

1.1	The RSC receives its authority from the President of Brock University.  It is responsible for coordinating and controlling all activities related to radiation safety.

1.2	The RSC consists of:

1.2.1	A person with competence in radiation safety and practical knowledge of the nature and use of radioactive material, who will act as the Radiation Safety Officer (RSO);

1.2.2	A representative of each Faculty using radioactive material, preferably an individual with the knowledge and training in the use of radiation emitting sources and materials, nominated by the Dean of the Faculty;

1.2.3	A representative of Human Resources, Environment, Health and Safety of an individual with appropriate knowledge and experience in radiation safety, nominated by the Executive Director;

1.2.4	An observer from the Brock University Faculty Association;

1.2.5	One staff member from Technical Services with appropriate knowledge and expertise in radiation emitting sources and materials and radiation safety;

1.2.6	The Associate Vice-President (AVP) Research and International Development who shall serve as Chair.

1.3	The RSC shall meet at least two times per calendar year and more frequently as required or at the call of the Chair;

1.4	To fulfill its responsibilities, the RSC has the following authority:

1.4.1	To review applications (Appendix 8) for the use of and to authorize and control by the issue of internal permits the use of radiation emitting sources and materials within the limits of the radioisotope licenses issued by the Canadian Nuclear Safety Commission (CNSC).  Internal permit holders are bound by the regulations stated in section A.3.

1.4.2	To administer and retain a record of training provided through a radiation safety program to ensure that all persons involved in the handling of radioisotopes have appropriate training to enable them to perform their duties safely.  The program is designed to ensure that radiation exposure by any person be kept as low as reasonably achievable (ALARA principle).

1.4.3	To suspend the use, at the University, of radiation emitting sources and materials, regardless of any other source of authorization.

1.4.4	To inform the Associate Vice President, Research and the Deans of all Faculties in which radioactive sources are used of the hazards associated with all radiation emitting sources and materials.

1.4.5	To produce and continually review the Brock University Radiation Safety Policies.

1.4.6	To maintain a file on all radiation laboratories and operations.  The file shall include a copy of internal permit applications and authorized permits, an inventory of radioactive sources, a copy of RSC laboratory inspection reports, a copy of Personnel dosimeter records.  Such files shall be considered active until all of the radioactive sources have been accounted for, either by disposal or by safe storage.

1.4.7	To approve the purchase of radioactive materials and x-ray sources in accordance with the specifications of authorized projects.

1.4.8	To approve the storage of radioactive waste on campus and to organize the removal of radioactive waste from the campus.

1.4.9	To administer a personnel dosimetry service on behalf of the users and to maintain all necessary records.

1.4.10 To receive information from the Office of Research Services of research applications that may require oversight of the RSC and to inform the Office of Research Services of researchers who require a radioisotope permit in order to carry out their proposed research programs and/or operate research equipment.

1.4.11	Annual Reports of the RSC will be filed with the Office of Associate Vice-President Research.

1.4.12	To notify permit holders and radiation users of RSC meetings and to post agenda and minutes of meetings for non-in camera sessions.

2.     Radiation Safety Officer (RSO)

2.1 The responsibilities of the RSO are:

ADMINISTRATION

2.1.1	To act on behalf of the RSC in the day-to-day administration of these policies and procedures.

    - To serve as a resource to the RSC on matters related to these policies and procedures and on institutional compliance with CNSC regulations.

INSPECTION and MONITORING

    - With the assistance of the RSC, to periodically inspect and approve according to Government standards, all proposed radiation laboratories and to inspect annually all such laboratories.

2.1.4	With the assistance of the RSC, to supervise the leak testing of all large sealed radioisotope sources as required by Government regulations and make periodic checks on contamination in radioisotope laboratories.

RECORD KEEPING

2.1.5	To ensure that each incident involving radioisotope spillage or exposure to ionizing radiation is evaluated to decide if decontamination procedures and/or medical examination should be carried out.

    - To receive documentation for internal permit applications and to administrate their authorization.

    - To ascertain the requirements for personnel monitoring and approve the use of          radiation detection instruments and dosimeters.

    - To maintain a record of the training provided to permit holders, radiation users,

and other staff who may be required to handle such materials or provide

maintenance, security or related services for laboratories in which radiation

emitting sources and materials might be used or stored.

EDUCATION

2.1.9	To ensure that general advice on radiation hazards and protection is available to permit holders, users.

    - To provide regular training to permit holders, radiation users, the RSC and other staff who may be required to handle such materials or provide maintenance, security or related services for laboratories in which radiation emitting sources and materials might be used or stored.

2.1.11  Where minors may be engaged in experiential education programs (e.g., Scientifically Yours or high school co-op placements) in laboratories using radiation emitting sources and materials, such individuals require special written permission of the RSO, education and training in the use of such materials and written school and/or parental permission prior participating in activities in such laboratories.

REPORTING

2.1.12	To present an annual report (Jan.-Dec.) of the activities of the RSC to the AVP Research who shall submit to the CNSC.

3.     Internal Permit Holders and Responsibilities

3.1	Will be faculty members with education, training and experience in the safe handling of radiation emitting sources and materials.  Each permit holder is responsible for:

3.1.1	Provision of adequate facilities, equipment, instruments, supervision and instruction to control radiation hazards and to comply with the University's radiation protection standards.

3.1.2	Maintain an up-to-date listing with the RSO of the rooms in which radioactive material is stored or handled, and of rooms in which radiation-emitting equipment is used.

3.1.3	Maintain an accurate inventory of radioactive materials used in his/her project, and to ensure that the project does not exceed its allotment of radioactive material.  An up-to-date inventory must be submitted annually at the request of the RSO, outlining the quantity received, disposed and the balance.

3.1.4	Keeping records of the disposal of radioactive material.

3.1.5	Ensuring that personnel wear appropriate film badges or pocket dosimeters, if required.

3.1.6	Allowing only authorized persons to enter rooms that are specified as restricted areas for reasons of radiation protection. These rooms must be kept locked when unattended. All radioactivity inventories must be kept in a secure location.

3.1.7	Posting of warning signs and labels as required by CNSC Regulations and Radiation Safety Policy.

3.1.8	Posting appropriate basic or intermediate laboratory rules in the radioisotope laboratories.   Copies may be obtained from the RSO.

3.1.9	Establishing a daily laboratory procedure to ensure that at the end of the workday:

(a) Measurements have established that external radiation and contamination levels are within permissible limits that are given in Appendix 5.  Records are to be kept and maintained for 3 years;

(b) Radiation sources are properly labeled and stored;

(c) Experiments that will be in progress after normal working hours will be properly attended;

(d) Each laboratory is secured against unauthorized access.

3.1.10	Performing leak tests capable of detecting the presence of radioactive material on each specified sealed source (normally sources less than 50 MBq and gaseous sources such as Krypton 85 are excepted) at least once every 6 months for sources in active use, 12 months for sources within a radiation device, 2 years for sources in storage and after any incident which could result in source damage. All test samples will be sent to a CNSC approved company for measurement. Records shall be maintained for at least 3 years.  If removable contamination in excess of the CNSC contamination criterion for the particular radioisotope is detected, the source shall be isolated, its use immediately discontinued, and the CNSC notified.

3.1.11	Report all radiation incidents in accordance with Radiation Incidents - Emergency Procedures, Section G.

3.1.12	Train and educate all radiation users in lab specific safe handling, hazards, and use of radiation emitting sources and materials. Radioisotope Use and Disposal information (Appendix 3) must be documented for each user and submitted to the RSO.

4.     Radiation User and Responsibilities

  - Each individual who may use radioactive materials or sources of ionizing radiation is responsible for complying with the procedures and precautions contained in the Radiation Safety Policy and Procedure and with those prescribed in the permit.

  - Every user must have received appropriate training and have been informed of the risks

associated with exposure to radiation prior to commencement of work procedures.

All users must receive lab-specific training and understand emergency lab procedures given

by the permit holder, or by an experienced user delegated by the permit holder to provide

such training. Each user must also attend the radiation safety awareness training course,

which is made available once every semester.  Interim training by their supervisor is

acceptable until a user can attend the formal training session given by the RSO.  Every 3 years all users shall attend a refresher General Safety Training session and satisfactorily complete the test.

4.2.1   A record of the training received by radiation users shall be maintained by the

permit holder for lab specific training and by the RSO for general education and

training.

# B.     RADIATION - WORK PROCEDURES

1.	Definition of a Nuclear Energy Worker

1.1	"Nuclear energy worker" means any person who in the course of his work, business or occupation, is likely to receive a dose of ionizing radiation in excess of any dose specified in Appendix 5, Table 2.  Minors (persons under the age of eighteen) may not be employed as "nuclear energy workers".

1.2	Each person who, in the opinion of the RSC, may be exposed to external or internal radiation (except prescribed medical treatment) in excess of the limits listed in Appendix 5, Table 2, will be considered as a "nuclear energy worker".

1.3	Each "nuclear energy worker" as defined in Section 1.1, must register with the Chair of

the RSC and receive a radiation protection instruction-interview with the RSO.

2.	Medical Examinations

2.1	Each person registered as a "nuclear energy worker" shall be required to have medical examinations of such a nature and at such intervals as required on the advice of the CNSC.

2.2	Any person using radioactive substances may be removed from such work on the advice of the CNSC.

3.	Exemption Quantity of a Radioisotope

3.1	"Exemption Quantity" means, in relation to an isotope, the quantity listed in the CNSC document “Nuclear Substances and Radiation Devices Regulations”.  Sources less than exemption quantities are license exempt according to the Regulations.  However, on the Brock University campus the exemption quantity is used only as a reference and approval must be obtained for work involving any radioactive material, even quantities less than the exemption quantity.  Some exemption quantities are listed in Appendix 7.

4.	Permissible Exposures

4.1	Exposure to ionizing radiation shall be kept as low as reasonably achievable (ALARA principle).

4.2	The exposure from sources of radiation shall normally be controlled in such a way as to provide assurance that no individual or user, except a nuclear energy worker, shall receive an absorbed dose in excess of the values listed in Appendix 5, Table 2.  Non-nuclear energy workers should not incur a whole body dose in excess of 0.1 mSv/week.

4.3	Maximum permissible doses for nuclear energy workers are listed in Appendix 5, Table 2.  Nuclear energy workers shall not incur a whole body dose in excess of 1 mSv/week.

5.	Personal Monitoring

  - Each radiation user must ensure that radiation exposure and contamination are properly controlled during the work with radiation so that the prescribed permissible limits given in Appendix 5 are not exceeded.

  - Personal thermoluminescent dosimetry badges (TLD badges) to users who perform work that meets or exceeds the criteria given in Appendix 5, Table 3. All users must have received radiation safety training prior to being issued a badge and prior to starting work procedures involving radioisotopes.

5.3	Bioassay Program for Iodine-125 and Iodine-131 Users

All users of the radioisotopes I-125 or I-131 who perform work that exceeds the limits in the following table, within a 24 hour period, must undergo thyroid screening within FIVE days.

Work Site		Quantity Exceeded Within a 24 hour Use

Open room		                 2     MBq

Fume hood		               200   MBq

Glove box		            20,000 MBq

All users who plan on exceeding these limits must obtain written permission from the Radiation Safety Committee prior to use so that bioassays can be scheduled.

Bioassays will be arranged with the Niagara Health System. Results will be returned to the RSO who will notify the user of the outcome.

In addition :

- any individual that is involved in a spill of greater than 2 MBq of Iodine-125 or Iodine 131

or

- any individual upon who an external contamination of Iodine-125 or Iodine-131 is detected

must undergo thyroid screening within FIVE days as well.

6.     Laboratory Monitoring

  - Permit holders and all users of radioisotopes must monitor radioactive contamination on work surfaces as part of each work procedure. Design plans for areas where radioisotopes are used should indicate the areas tested for contamination. Wipe test records must be stored in the laboratory and must correspond to inventory records indicating the date of use.

  - Wipe test procedures. Cut a 5 cm x 1 cm piece of absorbent filter paper and moisten it with alcohol or water. Rub the last 1 cm over the test surface to cover a total area of approximately 100 cm. Allow the paper to dry and cut the final, exposed 1 cm x 1 cm piece into a scintillation vial. Add 5ml of scintillation fluid, cap the vial and count for 1 –2 minutes in a scintillation counter with a broad spectrum (wide open) channel setting. Write the results in the lab record book and staple the counter printer output to the page. Decontaminate any controlled work area with a recorded measurement of 300 Bq/cm2 or greater or any public area with a count of 30 Bq/cm2 or greater for any Class C substances, 30 Bq/cm2 or greater or any public area with a count of 3 Bq/cm2 or greater for any Class B substances or 3 Bq/cm2 or greater or any public area with a count of 0.3 Bq/cm2 or greater for any Class A substances (see Appendix 7, Table 2).  The decontaminated area must be re-tested to ensure that contamination does not exceed the acceptable limits above.

  - Contamination meter procedures. All survey meters must be checked for efficiency at least annually. To do this check, obtain three certified radioisotope standards that emit beta particles with maximum energies from 50 KeV to 2 MeV. Place one of the standards on a flat surface and hold the detector probe over the source at a distance of 1 cm. Record the count rate. Repeat with the other standards and calculate efficiency as:

Efficiency =   Counts from standard – background counts

Known standard activity

Graph efficiency vs. beta-particle maximum energy.

For daily operation, check the battery of the instrument. Hold the detector approximately 1 cm from the surface to be tested and move the detector slowly over the surface. Do not touch the surface. Record the measurement in the lab record book. Surface contamination is calculated by:

Surface contamination =   Counts – background	x			1

Probe area (cm2)	Efficiency

Decontaminate any controlled work area with a calculated measurement of 300 Bq/cm2 or greater or any public area with a count of 30 Bq/cm2 or greater for any Class C substances, 30 Bq/cm2 or greater or any public area with a count of 3 Bq/cm2 or greater for any Class B substances or 3 Bq/cm2 or greater or any public area with a count of 0.3 Bq/cm2 or greater for any Class A substances (see Appendix 7, Table 2). The decontaminated area must be re-tested to ensure that contamination does not exceed the acceptable limits above.

7.	Decontamination Procedures

Good work procedures will help to prevent contamination problems. If contamination does occur, use one or more of the methods given below. The initial method of choice is scrubbing with minimal amounts of hot water with decontamination solutions. All waste materials must be disposed of appropriately. Consult the RSC for large-scale problems.

Method	Application		Advantages 			Disadvantages

Water and	All surfaces		Very effective if done	Poor for porous

Detergent	immediately		surfaces

after spill

Decon		Nonporous		Contamination remains	Long soaking

surfaces		in solution; non-toxic	required; little

penetration

Organic	Nonporous		Quick				May be toxic or

Solvents	surfaces						flammable.

Abrasion	Nonporous		Remaining contamination	Ineffective on

surfaces		very low			porous surfaces

8.	Authorization for Use of Radioisotopes

8.1	Prior to possessing or using radioactive material, authorization must be obtained from the RSC.

8.1.1	Complete an "Application for Radioisotope Permit" and forward to the RSO.  A copy of the application is illustrated in Appendix 8.  The requested "possession quantity" of radioactive material relative to the exemption quantity, Appendix 7 Table 1, and the limits given in Appendix 6, determines the manner in which approval is given.

(a)  	for quantities less than 100 exemption quantities, immediate conditional approval may be given.

(b)	for quantities less than the limits given in Appendix 6, but greater than 100 exemption quantities, the application must be circulated to all members of the RSC and if no objections are received, approval may be granted seven days after circulation.

(c)	for quantities greater than the limits given in Appendix 6, the application must be considered at a meeting of RSC.

8.1.2	After approval, an internal permit will be issued and a copy of the application will be returned to the applicant stating the conditions of approval.

8.1.3	If any changes are to be made to a Radioisotope operation, the permit holder must apply in writing to the RSO for an amendment to the permit.

8.2	Each room or laboratory in which radioactive material is to be handled or stored must be approved for use by the RSC.  A copy of the internal permit must be posted in plain view in the room.

9.	Procurement of Radioisotopes

  - To receive authorization for radioisotope purchase, all permit holders must complete and submit their radioisotope purchase requisition. Authorization to sign radioisotope purchase orders is vested in designated members of the RSC.

    - Personnel receiving radioisotope shipments from off campus must be authorized by the RSC. Only those individuals who have had Transport of Dangerous Goods training will receive this authorization. Shipments must be inspected in the receiving laboratory as soon as possible for damage or contamination of the packaging material. CNSC poster INFO-0426 “Receiving of Radioactive Packages” (see Appendix 9) should be used as a guide. The RSC must be notified of any damaged materials.

10. 	Use and Storage of Radioisotopes

10.1	Any laboratory in which radioactive material is used or stored shall be classified as a radioisotope laboratory and authorized by the Brock University RSC.

10.2	The quantity of radioactive material used or stored in a laboratory shall not exceed that specified on the internal radioisotope permit issued by RSC.  The permissible radioactivity, and appropriate modifying factors, that can be accommodated in a normal radioisotope laboratory may be estimated from the data in appendices 5 and 6.

10.3	Radioactive material shall be kept or stored in a manner that:

10.3.1	the storage container is labeled in accordance with the regulations of Radiation - Signs and Labels, Section E.

10.3.2	provides adequate radiation shielding as prescribed in the user’s permit.

10.3.3	provides adequate protection against theft, fire, explosion, and flooding or accidental breakage of primary storage containers.

10.4	Containment and disposal of radioactive waste will be in accordance with the procedure outlined in Radioisotopes – Safe Handling, Waste Storage, and Disposal, Section C.

11.	Termination of Radioisotope Operation: Decommissioning Rooms

11.1	The permit holder is responsible for removal of all radiation-warning signs, monitoring all areas, decontaminating areas to meet CNSC limits, updating and maintaining all records for a period of three years after the termination of operations. The RSC must be notified of all users who were issued personal dosimeters and whose names will be removed from the dosimetry list. Complete Section 3 of the Application for Radioisotope Permit and forward to the RSO.  Account for all radioisotopes acquired and used for the project in the preceding year.  Indicate the amounts and storage location of any remaining radioactive material.

11.2	Radioactive material may not be transferred to the possession of any person nor used for any purpose in any place other than originally authorized without prior written approval of the RSC.

  - The RSC will arrange and coordinate the disposal and/or relocation of radioisotopes.

  - Residual contamination in decommissioned rooms must not exceed 30 Bq/cm2 in any area.

# C.     RADIOISOTOPE – Safe Handling, Waste Storage and Disposal

1.    Responsibility of Permit Holder in Handling and Disposal

Each radioisotope permit holder is responsible for the control, containment and identification of active wastes generated within his or her laboratory, and removal to waste containers provided by the RSC.  They must sign a waste disposal form each time radioactive materials are used and disposed.  They must also ensure the RSO is informed immediately if a user will no longer be working under their permit.

2.    Responsibility of Radioisotope Users in Handling and Disposal

All users must keep accurate records of their use and disposal of radioisotopes.  If a user will no longer be working with radioisotopes, he or she must ensure all radioisotope use and disposal has been properly documented.  The user must inform the RSO when they are completed, so their name can be removed from the list of users on the internal operating permit.  The TLD badge of the user must be returned to the RSO at the completion of their work.

3.	Containment

3.1	In general, laboratory radioactive wastes should be segregated and contained such that:

3.1.1	the total amount of radioactive material put in any container must be controlled so that the radiation intensity at 30 centimeters is less than 0.5 mSv/hr.

3.1.2	material must not be put into waste collection containers if there is any possibility of a chemical reaction.

3.2	A record must be kept of the quantity and kinds of radioactive material disposed into each collection container.

4.	Disposal of Level One Radioactive Wastes

4.1	Radioactive wastes may be disposed into the environment if the level of radioactivity is low enough to satisfy the following criteria:

4.1.1	Via the sewer, annual disposal limits for the University as a whole of some radioisotopes in a form readily soluble and dispersible in water are:

H-3	    1000000MBq/yr

C-14	        10000 MBq/yr

P-32	                1 MBq/yr

S-35	          1000 MBq/yr

I-125	            100 MBq/yr

I-131	             10 MBq/yr

4.1.2	To the atmosphere provided that the resulting concentration in air at the point of release is less than:

H-3	    37KBq/m3

I-125	 0.03 KBq/m3

I-131      0.175 KBq/m3

4.1.3   Via normal garbage provided that the activity concentration is less 				                     than:

H-3	    37MBq/kg

C-14	   3.7 MBq/kg

P-32	 0.37 MBq/kg

S-35	 0.37 MBq/kg

I-125	0.037 MBq/kg

I-131	0.037 MBq/kg

5.	Disposal of Level Two Radioactive Wastes

5.1	Solid Wastes

5.1.1	Active dry waste should be collected in an appropriately labeled container lined with a heavy polyethylene bag.  Sharp objects should first be placed in protective containers.  Active dry waste should be removed to waste containers provided by RSC.

5.1.2	Disposal of other active waste (carcasses, sealed sources) must be done in consultation with RSC.

5.2	Liquid Waste

5.2.1	Non-aqueous radioactive waste solutions, including scintillation counter fluids, should be emptied into the appropriately designated containers.  An inventory of volumes, chemical nature of all materials, and maximum level of radioactivity must be kept on the form adjacent to the container.  No halogenated compounds are to be disposed of in this way.

5.2.2	Aqueous radioactive liquids should be emptied into an appropriately designated container.

5.2.3	Radioactive liquid wastes are stored in the flammable stores room and their disposal should be arranged through RSC.

# D.     Radioisotopes- General Requirements and Precautions

Observe Radioisotope Laboratory Rules posted in the Laboratory.

# E.     Radiation Signs and Labels

1.	 Radioisotope Laboratories

1.1	Laboratories in which radioactive isotopes are present shall have posted at each entrance to that laboratory, a sign with the radiation warning symbol and the words "Caution - Radioactive Materials".  This sign must also include the names and local and home telephone number of responsible persons to be contacted in an emergency, and other special emergency instructions.

1.2	Cupboards, cabinets, refrigerators and other containers used to store active materials in the radioisotope laboratory must be identified with a radiation-warning label.

1.3	Primary storage containers must be identified with a radiation-warning symbol and information respecting the nature, form, quantity and date of measurement of the radioactive isotopes contained within.  Labelling is not required for containers such as beakers, flasks and test tubes used while the responsible individual is present.

2.	Other Radiation Areas

  - Areas, rooms, or enclosures where the activity of any radioisotope may exceed 100 exemption quantities or where radiation fields may exceed 25 Sv/hr must be posted with a radiation warning sign with information concerning the nature and extent of the radiation hazard.  Appropriate radiation warning labels are required on all x-ray and other ionizing radiation producing devices.

3.     Unauthorized Use of Radiation-Warning Signs

3.1	Radiation-Warning Signs must not be used in any location where radioisotopes are not present.

# F.	 X-ray And other Radiation Producing Devices - Work Procedures

(a)	A copy of the "Regulation made under the Public Health Act respecting X-ray Safety" shall be available in each area where X-ray equipment is used.

(b)	Care shall be taken to avoid personal exposure to a primary X-ray beam.  Exposure in areas where radiation scatter is possible shall be kept to a minimum.

(c)	No interlock or other safety device shall be deliberately defeated.  No "live" adjustments or alignments shall be made with any safety cover removed, except where the procedure has been specifically checked for safety and the method is documented.

(d)	When a modification to an X-ray source is introduced, either temporary or permanent, in operating technique, equipment arrangement or in ancillary equipment the modification shall be checked for safety and all changes documented.

(e)	Any person to whom a Monitoring Film Badge has been issued must wear it when in the vicinity of an X-ray source.

(f)	Any defect in X-ray equipment resulting in possible radiation hazard shall be reported to the responsible user who will inform the RSO.

(g)	Precautions must be taken to prevent personal injury from laser beams, in particular special care must be taken to ensure that light from a laser does not enter the eye of any person.

(h)	Eye protection should be worn when approaching an Ultra Violet light source.

# G.     Radiation Incidents - Emergency Procedures

1.	All radiation incidents (exposure, spill, contamination, loss of source, etc.) must be reported immediately to the RSO. The RSO must report any major incidents to the CNSC within 24 hours and a written follow-up report must be sent to the CNSC within 21 days.

2.	Persons exposed to internal or external radiation present no hazards to others.  Such persons will be placed under medical supervision.  Re-entry into the radiation area is prohibited, until authorized by the RSO or any member of the RSC.

3.	In the event of personnel contamination, clothing and other articles that may be contaminated must be left in the controlled area.  Contamination should be washed from the skin under running water until the arrival of experienced assistance.

4.	In the event of a large spill, laboratory contamination, or environmental contamination, vacate the area.  Do not attempt decontamination procedures and do not remove contaminated articles from the area.  The RSO or a member of RSC will direct decontamination operations and authorize re-entry into the radiation area. Procedures used to decontaminate the area will be as outlined in Section B 7: “Decontamination Procedures” of the Radiation Safety Manual.

APPENDIX 1: Emergency Procedures Flowchart

RADIOACTIVE SPILLS RESPONSE

INFORM NEARBY INDIVIDUALS OF SPILL

EVACUATE AND SECURE SPILL AREA,

BUT HOLD AFFECTED INDIVIDUALS NEARBY

ARE THERE

MAJOR INJURIES?

IF THERE IS ANY CHANCE THAT THE VICTIM HAS BEEN DIRECTLY CONTAMINATED, INFORM MEDICAL RESPONDERS ABOUT THE NATURE OF SPILL MATERIAL (RADIOISOTOPE, LIQUID or SOLID, MAXIMUM ACTIVITY, …)

EXPOSURE TO BODY?

YES

ADMINISTER ANY MINOR FIRST AID IF NECESSARY

YES

NO

MONITOR PERSON(S) FOR POSSIBLE CONTAMINATION

REMOVE CONTAMINATED CLOTHES

PERSONAL DECONTAMINATION PROCEDURE

(see procedure on reverse side)

HAVE AN UNCONTAMINATED PERSON CALL FROM A SAFE LOCATION :

RADIATION SAFETY OFFICER at ext. 3417 (day)

CAMPUS SECURITY at ext. 3200 (night)

IN ALL CASES :

1.     NOTIFY THE PERMIT HOLDER

MONITOR AREA (WIPE TESTS) & DOCUMENT RESULTS

MONITOR PERSONNEL

REPORT ANY EXTERNAL RELEASE TO RADIATION SAFETY OFFICER

FILL OUT A BROCK INCIDENT/INJURY FORM WITH YOUR SUPERVISOR & SUBMIT TO EH&S WITHIN 24 HRS

SURFACE DECONTAMINATION

PROCEDURE

(see procedure on reverse side)

RE-MONITOR CONTAMINATED PERSON(S)

RE-MONITOR CONTAMINATED PERSON(S)

IS

CONTAMINATION AT BACKGROUND LEVELS?

NO

YES

NO

CALL 911 AND CALL CAMPUS SECURITY at ext. 3200

SECURITY WILL DIRECT THE AMBULANCE TO THE CORRECT LOCATION

Approved by: Radiation Safety Committee

Approval Date: July 2007

Personal decontamination : - Non-affected individuals attempting to aid in decontamination must put on protective gloves and a lab coat. - Using a survey meter, attempt to localize the contaminated area on the affected person’s body. - If eye is exposed- rinse eyes thoroughly at an eye wash station for at least 5 minutes. If ingested- rinse mouth out with water. If the skin is unbroken – using warm water and a mild non-abrasive soap, lather the area and rinse, taking care to not widen the contaminated area. If the skin is broken – using warm water carefully dab the affected area and gently wipe away from the wound, so as not to spread the radioactive material into the open wound. Avoid widening the contaminated area. If possible, promote slight bleeding so as to flush any radioactive material out of the wound. - Place all contaminated clothing and clean-up material into plastic bags, seal and label as radioactive waste, as well, include any known information (ie. radioisotope, maximum activity of spill, etc.) - Seek medical attention, where appropriate.

Surface decontamination : 1) Put on protective gloves, disposable footware and a lab coat. 2) Mark the area of spill with tape, marker or a wax pencil. - If the spill is: wet - place absorbent material (paper or granular) on the spill dry - use moist paper - Wash surface with detergent and hot water, using minimal amounts of water. - Wipe towards the centre of the spill, being careful not to spread the contamination outside the marked area. - Place all contaminated clean-up material into plastic bags, seal and label as radioactive waste, as well, include any known information (ie. radioisotope, maximum activity of spill, etc.). - Wipe test the area. - Using a survey meter, monitor decontamination personnel. - If repeated washing with detergent and hot water are not sufficient to reduce the contamination level below 300 Bq/cm2 for a controlled work area, or 30 Bq/cm2 for any public area, the following methods can be employed on non-porous surfaces: Method; Advantages; Disadvantages. Decon; Contamination remains in solution (non-toxic); Long soaking required, little penetration. Organic Solvents; Quick; May be toxic or flammable. Abrasion; Remaining contamination very low; Ineffective on porous surfaces.

APPENDIX 2: Inspection Report of Radioisotope Facilities

Permit holder:						Date:

Site:

Items				Compliance ()		Comments

Security, door

Security, sources

Up-to-date permit

Up-to-date user list

Food or drink in lab

Localization of procedures

Shielding

Inventory

Use records

Wipe tests

Waste disposal records

Disposal of scintillation

vials:

Other contamination

monitoring:				 (Specify:				)

Other:					 (Specify:				)

Signature:			_____________________________________

Frank Benko, Radiation Safety Officer

APPENDIX 3: Radioisotope Use And Disposal

ATTACH THE PRODUCT INFORMATION SHEET TO THIS RECORD

USE A SEPARATE FORM FOR EACH SHIPMENT

ISOTOPE:

AMOUNT (µCi and µL):

LOT NUMBER/UNIQUE IDENTIFIER:

REFERENCE DATE (D/M/Y):

| USE DATE (D/M/Y) | USER NAME | AMOUNT USED | AMOUNT REMAINING | DISPOSAL: AMOUNT, DATE AND METHOD OF DISPOSAL | SUPERVISOR SIGNATURE* |
| --- | --- | --- | --- | --- | --- |

* The supervisor or designated person-in-charge must review and approve all entries for the use and disposal of radioisotopes.  Under no circumstances may a user dispose of radioactive materials without completing this document.  It is important that anyone who will no longer be working as a radioisotope user must complete all disposal records and inform the RSO.

APPENDIX 4: Lab Specific Training In Radioisotope Use

Permit holder:						Date:

Trainee:

Radiation Safety Topics (check if covered)

Apparel (gloves, lab coat):	         	______	Localization of procedures:

General safety procedures:	            	           Technique specific methods:

Waste disposal, solids:	            	           Waste disposal, liquids:

Wipe tests:			            	           Scintillation counting:

Disposal of scintillation

vials:

Other contamination

monitoring:			            	 (Specify:				)

Record keeping:

Room and isotope security:

Other:				            	 (Specify:				)

Signatures

Permit holder:

I have read the following:	⁪ Radiation Safety Manual

⁪ Brock’s Occupational Health and Safety Policy

⁪ Brock’s Working Alone Policy

I have attended: ⁪ Lab-specific training in radioisotope use and safe handling

⁪ Radiation safety awareness training by the RSO

I acknowledge my responsibilities as a user, as outlined in the Radiation Safety Manual.

Trainee:

(signature)

Send completed form to Frank Benko, Radiation Safety Officer

APPENDIX 5: Maximum Permissible Levels

Dosage Units

"The seivert (Sv)" means a dose of ionizing radiation that has the same biological effects as 200-250 kilovolt X-rays whose energy is absorbed by the body or any tissue or organ thereof in amount 1 joules per kilogram.

For purposes of evaluating the biological effects of different radiations "Weighting Factors" are employed as given in Table I.

TABLE 1

| Radiation | Weighting Factor |
| --- | --- |
| 200-250 KeV X-rays | 1 |
| -rays, - rays and electrons | 1 |
| Thermal neutrons | 3 |
| Fast neutrons (greater than 0.8 MeV), protons and -particles | 10 |
| Heavy recoil nuclei | 20 |

Maximum Permissible Doses

TABLE 2

|  | Members of the Public | Nuclear Energy Workers |
| --- | --- | --- |
| Organ, Tissue | mSv per year | mSv per year |
| Whole body | 1 | 50* |
| Skin | 50 | 500 |
| Extremities: hands and feet | 50 | 500 |
| Extremities: head, eye lens | 15 | 150 |

*	If the person is known to be pregnant, the dose shall not exceed 4 mSv during the period of the pregnancy.

APPENDIX 5 (Cont'd)

Personal Dosimetry

TABLE 3

| Radioisotope | Process or Quantity per Single Use | Whole Body TLD | Extremity: Wrist or Ring |
| --- | --- | --- | --- |
| H-3, C-14, S-35 | Not applicable. | Not required | Not required |
| P-32 | <50 MBq >50 MBq | Yes Yes | No Yes (Ring) |
| I-125, I-131 | Iodination RIA | Yes Yes | Yes No |

Maximum permissible contamination of work surfaces

The internal permit holder shall ensure that the levels of loose beta and gamma radioactive contamination on all normally accessible working surfaces in a radioisotope laboratory do not exceed 300 Bq/cm2 or greater or any public area with a count of 30 Bq/cm2 or greater for any Class C substances, 30 Bq/cm2 or greater or any public area with a count of 3 Bq/cm2 or greater for any Class B substances or 3 Bq/cm2 or greater or any public area with a count of 0.3 Bq/cm2 or greater for any Class A substances (see Appendix 7, Table 2). The contamination level may be averaged over an area not exceeding 100 square centimeters.

The maximum permissible levels of loose alpha contamination should be one tenth those specified for beta or gamma emitting isotopes.  For fixed contamination, maximum permissible levels should depend upon the measured radiation fields.  Notwithstanding these levels, licensees are expected to keep radiation fields and contamination levels as low as reasonably achievable.

APPENDIX 6: Permissible Activity In Radioisotope Laboratory

TABLE 1

1 Bequerel (Bq) = 1 disintegration per second (dps)

1 curie (Ci) = 3.7x1010 dps = 37 GBq.

| RADIOTOXICITY OF THE INDIVIDUAL RADIONUCLIDES | PERMISSIBLE LEVEL OF ACTIVITY |
| --- | --- |
| Very High | 100 Ci |
| High | 10 mCi |
| Moderate | ** 100 mCi |
| Slight | 1 Ci |

** Except for Mo/99mTc generators and 99Tc eluate, for which the permissible activity is 3 Ci.

TABLE 2

| TYPE OF OPERATION | MODIFYING FACTORS |
| --- | --- |
| Storage | x10 |
| normal chemical operations | x1 |
| dry and dusty operations | x .01 |

APPENDIX 7: Classification and Exemption quantity of Radioisotopes

TABLE 1

Exemption Quantity of Some Radioisotopes

- EXEMPTION QUANTITIES

| Column 1 | Column 2 | Column 3 |
| --- | --- | --- |
| Radioactive Nuclear Substance | Activity Concentration | Activity |
|  | (Bq/g) | (Bq) |
| Actinium 227 | 1 × 10–1 | 1 × 103 |
| Actinium 228 | 1 × 101 | 1 × 106 |
| Americium 241 | 1 × 100 | 1 × 104 |
| Americium 242 | 1 × 103 | 1 × 106 |
| Americium 242m | 1 × 100 | 1 × 104 |
| Americium 243 | 1 × 100 | 1 × 103 |
| Antimony 122 | 1 × 102 | 1 × 104 |
| Antimony 124 | 1 × 101 | 1 × 106 |
| Antimony 125 | 1 × 102 | 1 × 106 |
| Argon 37 | 1 × 106 | 1 × 108 |
| Argon 41 | 1 × 102 | 1 × 109 |
| Arsenic 73 | 1 × 103 | 1 × 107 |
| Arsenic 74 | 1 × 101 | 1 × 106 |
| Arsenic 76 | 1 × 102 | 1 × 105 |
| Arsenic 77 | 1 × 103 | 1 × 106 |
| Astatine 211 | 1 × 103 | 1 × 107 |
| Barium 131 | 1 × 102 | 1 × 106 |
| Barium 133 | 1 × 102 | 1 × 106 |
| Barium 140 | 1 × 101 | 1 × 105 |
| Berkelium 249 | 1 × 103 | 1 × 106 |
| Beryllium 7 | 1 × 103 | 1 × 107 |
| Bismuth 206 | 1 × 101 | 1 × 105 |
| Bismuth 207 | 1 × 101 | 1 × 106 |
| Bismuth 210 | 1 × 103 | 1 × 106 |
| Bismuth 212 | 1 × 101 | 1 × 105 |
| Bromine 82 | 1 × 101 | 1 × 106 |
| Cadmium 107 | 1 × 103 | 1 × 107 |
| Appendix 7 (Cont’d) | 1 × 104 | 1 × 106 |
| Cadmium 113m | 1 × 103 | 1 × 106 |
| Cadmium 115 | 1 × 102 | 1 × 106 |
| Cadmium 115m | 1 × 103 | 1 × 106 |
| Calcium 45 | 1 × 104 | 1 × 107 |
| Calcium 47 | 1 × 101 | 1 × 106 |
| Californium 246 | 1 × 103 | 1 × 106 |
| Californium 248 | 1 × 101 | 1 × 104 |
| Californium 249 | 1 × 100 | 1 × 103 |
| Californium 250 | 1 × 101 | 1 × 104 |
| Californium 251 | 1 × 100 | 1 × 103 |
| Californium 252 | 1 × 101 | 1 × 104 |
| Californium 253 | 1 × 102 | 1 × 105 |
| Californium 254 | 1 × 100 | 1 × 103 |
| Carbon 11 | 1 × 101 | 1 × 106 |
| Carbon 14 | 1 × 104 | 1 × 107 |
| Cerium 139 | 1 × 102 | 1 × 106 |
| Cerium 141 | 1 × 102 | 1 × 107 |
| Cerium 143 | 1 × 102 | 1 × 106 |
| Cerium 144 | 1 × 102 | 1 × 105 |
| Cesium 129 | 1 × 102 | 1 × 105 |
| Cesium 131 | 1 × 103 | 1 × 106 |
| Cesium 132 | 1 × 101 | 1 × 105 |
| Cesium 134 | 1 × 101 | 1 × 104 |
| Cesium 134m | 1 × 103 | 1 × 105 |
| Cesium 135 | 1 × 104 | 1 × 107 |
| Cesium 136 | 1 × 101 | 1 × 105 |
| Cesium 137 | 1 × 101 | 1 × 104 |
| Cesium 138 | 1 × 101 | 1 × 104 |
| Chlorine 36 | 1 × 104 | 1 × 106 |
| Chlorine 38 | 1 × 101 | 1 × 105 |
| Chromium 49 | 1 × 101 | 1 × 106 |
| Chromium 51 | 1 × 103 | 1 × 107 |
| Cobalt 55 Appendix 7 (Cont’d) | 1 × 101 | 1 × 106 |
| Cobalt 56 | 1 × 101 | 1 × 105 |
| Cobalt 57 | 1 × 102 | 1 × 106 |
| Cobalt 58 | 1 × 101 | 1 × 106 |
| Cobalt 58m | 1 × 104 | 1 × 107 |
| Cobalt 60 | 1 × 101 | 1 × 105 |
| Cobalt 60m | 1 × 103 | 1 × 106 |
| Cobalt 61 | 1 × 102 | 1 × 106 |
| Cobalt 62m | 1 × 101 | 1 × 105 |
| Copper 60 | 1 × 101 | 1 × 105 |
| Copper 64 | 1 × 102 | 1 × 106 |
| Copper 67 | 1 × 102 | 1 × 106 |
| Curium 242 | 1 × 102 | 1 × 105 |
| Curium 243 | 1 × 100 | 1 × 104 |
| Curium 244 | 1 × 101 | 1 × 104 |
| Curium 245 | 1 × 100 | 1 × 103 |
| Curium 246 | 1 × 100 | 1 × 103 |
| Curium 247 | 1 × 100 | 1 × 104 |
| Curium 248 | 1 × 100 | 1 × 103 |
| Dysprosium 159 | 1 × 103 | 1 × 107 |
| Dysprosium 165 | 1 × 103 | 1 × 106 |
| Dysprosium 166 | 1 × 103 | 1 × 106 |
| Einsteinium 253 | 1 × 102 | 1 × 105 |
| Einsteinium 254 | 1 × 101 | 1 × 104 |
| Einsteinium 254m | 1 × 102 | 1 × 106 |
| Erbium 169 | 1 × 104 | 1 × 107 |
| Erbium 171 | 1 × 102 | 1 × 106 |
| Europium 152 | 1 × 101 | 1 × 106 |
| Europium 152m | 1 × 102 | 1 × 106 |
| Europium 154 | 1 × 101 | 1 × 106 |
| Europium 155 | 1 × 102 | 1 × 107 |
| Fermium 254 | 1 × 104 | 1 × 107 |
| Fermium 255 | 1 × 103 | 1 × 106 |
| Fluorine 18 Appendix 7 (Cont’d) | 1 × 101 | 1 × 106 |
| Gadolinium 153 | 1 × 102 | 1 × 107 |
| Gadolinium 159 | 1 × 103 | 1 × 106 |
| Gallium 67 | 1 × 102 | 1 × 106 |
| Gallium 72 | 1 × 101 | 1 × 105 |
| Germanium 68 | 1 × 101 | 1 × 105 |
| Germanium 71 | 1 × 104 | 1 × 108 |
| Gold 195 | 1 × 102 | 1 × 107 |
| Gold 198 | 1 × 102 | 1 × 106 |
| Gold 199 | 1 × 102 | 1 × 106 |
| Hafnium 181 | 1 × 101 | 1 × 106 |
| Holmium 166 | 1 × 103 | 1 × 105 |
| Hydrogen 3 | 1 × 106 | 1 × 109 |
| Indium 111 | 1 × 102 | 1 × 106 |
| Indium 113m | 1 × 102 | 1 × 106 |
| Indium 114m | 1 × 102 | 1 × 106 |
| Indium 115 | 1 × 103 | 1 × 105 |
| Indium 115m | 1 × 102 | 1 × 106 |
| Iodine 123 | 1 × 102 | 1 × 107 |
| Iodine 125 | 1 × 103 | 1 × 106 |
| Iodine 126 | 1 × 102 | 1 × 106 |
| Iodine 129 | 1 × 102 | 1 × 105 |
| Iodine 130 | 1 × 101 | 1 × 106 |
| Iodine 131 | 1 × 102 | 1 × 106 |
| Iodine 132 | 1 × 101 | 1 × 105 |
| Iodine 133 | 1 × 101 | 1 × 106 |
| Iodine 134 | 1 × 101 | 1 × 105 |
| Iodine 135 | 1 × 101 | 1 × 106 |
| Iridium 190 | 1 × 101 | 1 × 106 |
| Iridium 192 | 1 × 101 | 1 × 104 |
| Iridium 194 | 1 × 102 | 1 × 105 |
| Iron 52 | 1 × 101 | 1 × 106 |
| Iron 55 | 1 × 104 | 1 × 106 |
| Iron 59 Appendix 7 (Cont’d) | 1 × 101 | 1 × 106 |
| Krypton 74 | 1 × 102 | 1 × 109 |
| Krypton 76 | 1 × 102 | 1 × 109 |
| Krypton 77 | 1 × 102 | 1 × 109 |
| Krypton 79 | 1 × 103 | 1 × 105 |
| Krypton 81 | 1 × 104 | 1 × 107 |
| Krypton 83m | 1 × 105 | 1 × 1012 |
| Krypton 85 | 1 × 105 | 1 × 104 |
| Krypton 85m | 1 × 103 | 1 × 1010 |
| Krypton 87 | 1 × 102 | 1 × 109 |
| Krypton 88 | 1 × 102 | 1 × 109 |
| Lanthanum 140 | 1 × 101 | 1 × 105 |
| Lead 203 | 1 × 102 | 1 × 106 |
| Lead 210 | 1 × 101 | 1 × 104 |
| Lead 212 | 1 × 101 | 1 × 105 |
| Lutetium 177 | 1 × 103 | 1 × 107 |
| Manganese 51 | 1 × 101 | 1 × 105 |
| Manganese 52 | 1 × 101 | 1 × 105 |
| Manganese 52m | 1 × 101 | 1 × 105 |
| Manganese 53 | 1 × 104 | 1 × 109 |
| Manganese 54 | 1 × 101 | 1 × 106 |
| Manganese 56 | 1 × 101 | 1 × 105 |
| Mercury 197 | 1 × 102 | 1 × 107 |
| Mercury 197m | 1 × 102 | 1 × 106 |
| Mercury 203 | 1 × 102 | 1 × 105 |
| Molybdenum 90 | 1 × 101 | 1 × 106 |
| Molybdenum 93 | 1 × 103 | 1 × 108 |
| Molybdenum 99 | 1 × 102 | 1 × 106 |
| Molybdenum 101 | 1 × 101 | 1 × 106 |
| Neodymium 147 | 1 × 102 | 1 × 106 |
| Neodymium 149 | 1 × 102 | 1 × 106 |
| Neptunium 237 | 1 × 100 | 1 × 103 |
| Neptunium 239 | 1 × 102 | 1 × 107 |
| Neptunium 240 Appendix 7 (Cont’d) | 1 × 101 | 1 × 106 |
| Nickel 59 | 1 × 104 | 1 × 108 |
| Nickel 63 | 1 × 105 | 1 × 108 |
| Nickel 65 | 1 × 101 | 1 × 106 |
| Niobium 93m | 1 × 104 | 1 × 107 |
| Niobium 94 | 1 × 101 | 1 × 106 |
| Niobium 95 | 1 × 101 | 1 × 106 |
| Niobium 97 | 1 × 101 | 1 × 106 |
| Niobium 98 | 1 × 101 | 1 × 105 |
| Nitrogen 13 | 1 × 102 | 1 × 109 |
| Osmium 185 | 1 × 101 | 1 × 106 |
| Osmium 191 | 1 × 102 | 1 × 107 |
| Osmium 191m | 1 × 103 | 1 × 107 |
| Osmium 193 | 1 × 102 | 1 × 106 |
| Oxygen 15 | 1 × 102 | 1 × 109 |
| Palladium 103 | 1 × 103 | 1 × 108 |
| Palladium 109 | 1 × 103 | 1 × 106 |
| Phosphorous 32 | 1 × 103 | 1 × 105 |
| Phosphorous 33 | 1 × 105 | 1 × 108 |
| Platinum 191 | 1 × 102 | 1 × 106 |
| Platinum 193m | 1 × 103 | 1 × 107 |
| Platinum 197 | 1 × 103 | 1 × 106 |
| Platinum 197m | 1 × 102 | 1 × 106 |
| Plutonium 234 | 1 × 102 | 1 × 107 |
| Plutonium 235 | 1 × 102 | 1 × 107 |
| Plutonium 236 | 1 × 101 | 1 × 104 |
| Plutonium 237 | 1 × 103 | 1 × 107 |
| Plutonium 238 | 1 × 100 | 1 × 104 |
| Plutonium 239 | 1 × 100 | 1 × 104 |
| Plutonium 240 | 1 × 100 | 1 × 103 |
| Plutonium 241 | 1 × 102 | 1 × 105 |
| Plutonium 242 | 1 × 100 | 1 × 104 |
| Plutonium 243 | 1 × 103 | 1 × 107 |
| Plutonium 244 Appendix 7 (Cont’d) | 1 × 100 | 1 × 104 |
| Polonium 203 | 1 × 101 | 1 × 106 |
| Polonium 205 | 1 × 101 | 1 × 106 |
| Polonium 207 | 1 × 101 | 1 × 106 |
| Polonium 210 | 1 × 101 | 1 × 104 |
| Potassium 40 | 1 × 102 | 1 × 106 |
| Potassium 42 | 1 × 102 | 1 × 106 |
| Potassium 43 | 1 × 101 | 1 × 106 |
| Praseodymium 142 | 1 × 102 | 1 × 105 |
| Praseodymium 143 | 1 × 104 | 1 × 106 |
| Promethium 147 | 1 × 104 | 1 × 107 |
| Promethium 149 | 1 × 103 | 1 × 106 |
| Protactinium 230 | 1 × 101 | 1 × 106 |
| Protactinium 231 | 1 × 100 | 1 × 103 |
| Protactinium 233 | 1 × 102 | 1 × 107 |
| Radium 223 | 1 × 102 | 1 × 105 |
| Radium 224 | 1 × 101 | 1 × 105 |
| Radium 225 | 1 × 102 | 1 × 105 |
| Radium 226 | 1 × 101 | 1 × 104 |
| Radium 227 | 1 × 102 | 1 × 106 |
| Radium 228 | 1 × 101 | 1 × 105 |
| Radon 220 | 1 × 104 | 1 × 107 |
| Radon 222 | 1 × 101 | 1 × 108 |
| Rhenium 186 | 1 × 103 | 1 × 106 |
| Rhenium 187 | 1 × 106 | 1 × 109 |
| Rhenium 188 | 1 × 102 | 1 × 105 |
| Rhodium 103m | 1 × 104 | 1 × 108 |
| Rhodium 105 | 1 × 102 | 1 × 107 |
| Rubidium 86 | 1 × 102 | 1 × 105 |
| Ruthenium 97 | 1 × 102 | 1 × 107 |
| Ruthenium 103 | 1 × 102 | 1 × 106 |
| Ruthenium 105 | 1 × 101 | 1 × 106 |
| Ruthenium 106 | 1 × 102 | 1 × 105 |
| Samarium 151 Appendix 7 (Cont’d) | 1 × 104 | 1 × 108 |
| Samarium 153 | 1 × 102 | 1 × 106 |
| Scandium 46 | 1 × 101 | 1 × 106 |
| Scandium 47 | 1 × 102 | 1 × 106 |
| Scandium 48 | 1 × 101 | 1 × 105 |
| Selenium 75 | 1 × 102 | 1 × 106 |
| Selenium 79 | 1 × 104 | 1 × 107 |
| Silicon 31 | 1 × 103 | 1 × 106 |
| Silver 105 | 1 × 102 | 1 × 106 |
| Silver 110m | 1 × 101 | 1 × 106 |
| Silver 111 | 1 × 103 | 1 × 106 |
| Sodium 22 | 1 × 101 | 1 × 106 |
| Sodium 24 | 1 × 101 | 1 × 105 |
| Strontium 85 | 1 × 102 | 1 × 106 |
| Strontium 85m | 1 × 102 | 1 × 107 |
| Strontium 87m | 1 × 102 | 1 × 106 |
| Strontium 89 | 1 × 103 | 1 × 106 |
| Strontium 90 | 1 × 102 | 1 × 104 |
| Strontium 91 | 1 × 101 | 1 × 105 |
| Strontium 92 | 1 × 101 | 1 × 106 |
| Sulphur 35 | 1 × 105 | 1 × 108 |
| Tantalum 182 | 1 × 101 | 1 × 104 |
| Technetium 96 | 1 × 101 | 1 × 106 |
| Technetium 96m | 1 × 103 | 1 × 107 |
| Technetium 97 | 1 × 103 | 1 × 108 |
| Technetium 97m | 1 × 103 | 1 × 107 |
| Technetium 99 | 1 × 104 | 1 × 107 |
| Technetium 99m | 1 × 102 | 1 × 107 |
| Tellurium 123m | 1 × 102 | 1 × 107 |
| Tellurium 125m | 1 × 103 | 1 × 107 |
| Tellurium 127 | 1 × 103 | 1 × 106 |
| Tellurium 127m | 1 × 103 | 1 × 107 |
| Tellurium 129 | 1 × 102 | 1 × 106 |
| Tellurium 129m Appendix 7 (Cont’d) | 1 × 103 | 1 × 106 |
| Tellurium 131 | 1 × 102 | 1 × 105 |
| Tellurium 131m | 1 × 101 | 1 × 106 |
| Tellurium 132 | 1 × 102 | 1 × 107 |
| Tellurium 133 | 1 × 101 | 1 × 105 |
| Tellurium 133m | 1 × 101 | 1 × 105 |
| Tellurium 134 | 1 × 101 | 1 × 106 |
| Terbium 160 | 1 × 101 | 1 × 106 |
| Thallium 200 | 1 × 101 | 1 × 106 |
| Thallium 201 | 1 × 102 | 1 × 106 |
| Thallium 202 | 1 × 102 | 1 × 106 |
| Thallium 204 | 1 × 104 | 1 × 104 |
| Thorium 226 | 1 × 103 | 1 × 107 |
| Thorium 227 | 1 × 101 | 1 × 104 |
| Thorium 228 | 1 × 100 | 1 × 104 |
| Thorium 229 | 1 × 100 | 1 × 103 |
| Thorium 230 | 1 × 100 | 1 × 104 |
| Thorium 231 | 1 × 103 | 1 × 107 |
| Thorium 232 | 1 × 101 | 1 × 104 |
| Thorium 234 | 1 × 103 | 1 × 105 |
| Thorium natural | 1 × 100 | 1 × 103 |
| Thulium 170 | 1 × 103 | 1 × 106 |
| Thulium 171 | 1 × 104 | 1 × 108 |
| Tin 113 | 1 × 103 | 1 × 107 |
| Tin 125 | 1 × 102 | 1 × 105 |
| Tungsten 181 | 1 × 103 | 1 × 107 |
| Tungsten 185 | 1 × 104 | 1 × 107 |
| Tungsten 187 | 1 × 102 | 1 × 106 |
| Uranium 230 | 1 × 101 | 1 × 105 |
| Uranium 231 | 1 × 102 | 1 × 107 |
| Uranium 232 | 1 × 100 | 1 × 103 |
| Uranium 233 | 1 × 101 | 1 × 104 |
| Uranium 234 | 1 × 101 | 1 × 104 |
| Uranium 235 Appendix 7 (Cont’d) | 1 × 101 | 1 × 104 |
| Uranium 236 | 1 × 101 | 1 × 104 |
| Uranium 237 | 1 × 102 | 1 × 106 |
| Uranium 238 | 1 × 101 | 1 × 104 |
| Uranium 239 | 1 × 102 | 1 × 106 |
| Uranium 240 | 1 × 103 | 1 × 107 |
| Uranium 240 | 1 × 101 | 1 × 106 |
| Uranium natural | 1 × 100 | 1 × 103 |
| Vanadium 48 | 1 × 101 | 1 × 105 |
| Xenon 123 | 1 × 102 | 1 × 109 |
| Xenon 129m | 1 × 103 | 1 × 104 |
| Xenon 131m | 1 × 104 | 1 × 104 |
| Xenon 133 | 1 × 103 | 1 × 104 |
| Xenon 135 | 1 × 103 | 1 × 1010 |
| Ytterbium 169 | 1 × 102 | 1 × 107 |
| Ytterbium 175 | 1 × 103 | 1 × 107 |
| Yttrium 90 | 1 × 103 | 1 × 105 |
| Yttrium 91 | 1 × 103 | 1 × 106 |
| Yttrium 91m | 1 × 102 | 1 × 106 |
| Yttrium 92 | 1 × 102 | 1 × 105 |
| Yttrium 93 | 1 × 102 | 1 × 105 |
| Zinc 65 | 1 × 101 | 1 × 106 |
| Zinc 69 | 1 × 104 | 1 × 106 |
| Zinc 69m | 1 × 102 | 1 × 106 |
| Zirconium 93 | 1 × 103 | 1 × 107 |
| Zirconium 95 | 1 × 101 | 1 × 106 |
| Zirconium 97 | 1 × 101 | 1 × 105 |

TABLE 2

Classification of Radioisotopes

The following radioisotopes have been grouped by the CNSC into class A, B or C based upon radiological properties. Contamination monitoring limits for public or work areas are dependent upon this classification. When using more than one radioisotope in a room, the radioisotope with the lowest limit must be used to determine the contamination limit, Class A, B or C that applies to that room. Contact the RSO for information on radioisotopes not included in this table.

| CLASS | RADIOISOTOPE |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Class A | All alpha emitters and their daughter isotopes. |  |  | Na-22 | Na-24 |
|  | Co-60 | Ir-192 | Sb-124 | Ta-182 | Zn-65 |
| Class B | As-74 | Au-198 | Br-82 | Co-58 | F-18 |
|  | Fe-59 | Ga-67 | Gd-153 | Hg-203 | I-131 |
|  | In-111 | In-114m | Nb-95 | Rb-84 | Rb-86 |
|  | Sc-46 | Se-75 | Sm-153 | Sn-113 | Sn-123 |
|  | Sr-85 | Sr-90 |  |  |  |
| Class C | Au-195m | C-14 | Ca-45 | Cd-109 | Ce-144 |
|  | Cl-36 | Co-57 | Cr-51 | H-3 | I-123 |
|  | I-125 | Ni-63 | P-32 | P-33 | Re-186 |
|  | Re-188 | Ru-103 | S-35 | Sr-89 | Tc-99 |
|  | Tc-99m | Tl-201 | Y-90 | Yb-169 |  |

APPENDIX 8: Operation and Termination Permits

1.	APPLICATION FOR RADIOISOTOPE PERMIT

Complete Section 1 and forward to the Radiation Safety Officer (RSO).  When approved, the application will be assigned a number and a copy will be returned to the permit holder.  The conditions of approval will be indicated in Section 2.

2.	TERMINATION OF RADIOISOTOPE LABORATORY OPERATIONS

Complete Section 3 and forward a copy, along with your records to the RSO.  Account for all radioisotopes acquired and used for your project.

Radioactive material may not be transferred to the possession of any person nor used for any purpose in any place other than originally authorized, without prior approval of the Radiation Safety Committee.

Radiation Safety Committee, Brock University

APPLICATION FOR RADIOISOTOPE PERMIT

SECTION 1

The completed application will be returned upon permit approval.

1.  APPLICANT (SUPERVISOR)

| NAME | DEPARTMENT | BUILDING | ROOM NO. | EXT. | EMAIL |
| --- | --- | --- | --- | --- | --- |
| ALTERNATE | DEPARTMENT | BUILDING | ROOM NO. | EXT. | EMAIL |

2.  TECHNICIANS and others handling the material

| NAME | DEPARTMENT | BUILDING | ROOM NO. | EXT. | EMAIL |
| --- | --- | --- | --- | --- | --- |

3.  BUILDING AND ROOMS where material will be handled

| DELIVERED TO | PREPARATION IN | USED IN | STORED IN |
| --- | --- | --- | --- |

4.  DESCRIPTION (Attach additional list if necessary)

Open Sources

|  | Activity |  |  |  |
| --- | --- | --- | --- | --- |
| Isotopes | Maximum Held At one Time | Per Experiment | Annual Requirement | Form |

Sealed Sources

| Isotopes | Activity | Manufacturer | Model No. | Type of Device |
| --- | --- | --- | --- | --- |

5.  STATEMENT OF INTENDED USE

6.   INSTRUMENTS TO BE USED FOR RADIATION DETECTION AND MONITORING

| TYPE | SENSITIVITY mR/hr |
| --- | --- |

7.    USE OF RADIOACTIVE MATERIALS WITH BIOHAZARDS

Provide a copy of your Biohazard permit, if applicable.

8.   WASTE DISPOSAL*

| ISOTOPE | PROPOSED METHOD OF DISPOSAL |
| --- | --- |

* Attach additional list if necessary

9.	ANTICIPATED DATE FOR PROJECT TERMINATION ______________________

10.	APPLICANT'S SIGNATURE  ____________________________

DATE ______________________

SECTION 2.	RADIOISOTOPE PERMIT

A.	Comments related to the application

B.	Conditions of approval

C.	Requirements:

1.	The proposed work shall be performed in the manner specified in Section 1 and 2B.  There shall be no change in the procedures, equipment, radioactive material, or places of use without prior approval of the Radiation Safety Committee.

2.	The use, storage and disposal of radioactive material shall be in conformity with,

(a)  The Canadian Nuclear Safety Commission regulations,

(b)  The provisions for radiation protection as stipulated in Brock University Radiation Safety Policy and Procedures.

APPROVED BY:  _________________________________________________

Radiation Safety Officer

Permit No.  ______________	Permit Holder    ________________________

Date approved			Termination Date

SECTION 3.	TERMINATION OF RADIOISOTOPE

LABORATORY OPERATIONS

Radioactive material may not be transferred to the possession of any person nor used for any purpose in any place other than originally authorized without prior approval of the Radiation Safety Committee.

A.	Date of Termination  __________________________

B.	Inventory*

| ISOTOPE | TOTAL ACQUIRED | DISPOSAL OR DECAY | TOTAL REMAINING |
| --- | --- | --- | --- |

*  Attach additional list if necessary, and disposal records

C. 	Proposal for Disposal or Storage of Remaining Isotopes

D.	Decontamination Test Results. Provide a floor plan of the lab to be decommissioned and indicate areas wipe tested. Include all wipe test records.

D.	Permit Holder's Signature  _______________________________

E.	Authorizing Signature  _____________________________

Radiation Safety Officer

APPENDIX 9: References

Brock Policies:

Occupational Health and Safety

www.brocku.ca/oehs/images/Occupational_Health_&_Safety_Policy.pdf

Working Alone

www.brocku.ca/oehs/graphics/Working_Alone_Policy.pdf

Canadian Nuclear Safety Commission

www.nuclearsafety.gc.ca/eng

Canadian Nuclear Safety Commission: Acts, Regulations, By-laws

www.nuclearsafety.gc.ca/eng/resource/regulations/

CNSC- ALARA Principle

www.nuclearsafety.gc.ca/pubs_catalogue/uploads/G129rev1_e.pdf

CNSC- Other Current Regulatory Documents

www.nuclearsafety.gc.ca/eng/resource/regulatory_docs/current_docs/index.cfm

INFO-0426 “Receiving of Radioactive Packages”

www.uottawa.ca/services/ehss/CNSC%20packages.pdf

Laser Institute of America:

www.laserinstitute.org

Ontario Health and Safety Act, Regulation 861, X-ray Safety

www.e-laws.gov.on.ca/DBLaws/Regs/English/900861_e.htm

Radiation Safety Institute of Canada:

www.radiationsafety.ca/Issues.htm